The Easidew Sampler  sampling system is  a self-contained  sensor/filter block with the

PRESSURE MEASUREMENT CONFIGURATION

(FACTORY SUPPLIED)

ability to measure either pressure or atmospheric dew points when used in conjunction with the Easidew sensor.

Connection Ports:

The entry and exit pipe connections are push fit types to accept PTFE, or similar, 6mm O/D pipe. A 0.5 metre length of PTFE is supplied which should be used as a pigtail from the OUTLET port, whether measuring in either the pressure or atmospheric mode.

145 Max. with pipe

85 Max. with pipe

Filter:

A 99.5% 0.3micron particulate filter cartridge is fitted as standard and is accessed via the filter cap. The condition of the filter should be checked at regular intervals to ensure optimum performance. Replacement cartridges can be ordered from Michell Instruments

(Part No. 96031). The filter cap is sealed with an O-ring, seated in the main body of the block. When inspecting or changing the filter, ensure the O-ring is correctly seated when retightening the filter cap.

Flow Control Valve:

A flow control valve is supplied factory fitted to the OUTLET port. This valve is design to set the optimum gas flow of between 1-5 litres per minute through the block.

The flow control valve should not be used as a flow shut-off valve.

Pressure Dew-Point Measurements:

The  Easidew  Sampler  system  is  factory  assembled  to  make  pressure  dew-point measurements. This is achieved by controlling the gas flow at the OUTLET port. The maximum operating pressure for the Easidew Sampler system is 10 barg.

Atmospheric Dew-Point measurements:

If  desired  the  block  can  be  easily  reconfigured  to  make  atmospheric  dew-point measurements by installing  the flow control valve in the INLET  port. Simply  swap positions of the flow control valve and the  gas  pipe connection coupling fitted at the inlet port. When reconfiguring the Easidew Sampler ensure that the system is leak tight before commencing measurements.

ALL MOUNTING HOLES M6

FILTER CAP

ATMOSPHERIC MEASUREMENT CONFIGURATION

COUPLING REMAINS IN PLACE FOR EITHER CONFIGURATION

OUTLET PORT

INLET PORT

OUTLET PORT

INLET PORT

Mounting:

The mounting of the Easidew Sampler system is not position critical. A mounting bracket is factory fitted. This bracket is easily removed and repositioned to provide a number of mounting options.

Alternatively it is possible to mount the Easidew Sampler system without the bracket by using the 2 off M6 x 5mm deep mounting fixings, pitch at 20mm, which are machined directly into the block.

CAUTION

Maximum operating pressure 10 Barg

It is important that the gas fittings are correctly tightened before use. Failure to do so will affect the pressure rating.

For use at higher pressures please contact Michell Instruments Ltd for details of alternative compression fittings.

Liquid Contamination

The Easidew Sampler does not provide protection from liquid contamination. If there is a possibility of liquid in the gas sample, additional filtration should be fitted before the Easidew Sampler.

Specification:

Gas Flow rate:	1-5 Nlmin-1

Operating Pressure:	10 bar (145psi)

Particulate Filter:	99.5 % removal of 0.3 microns

Environmental:	IP66 (NEMA 4)

Operating: -40ºC to +60ºC, 0-100 %rh condensing

Storage: -40ºC to +70ºC, 0-95 %rh non-condensing

Weight:	1.1Kg (1.3Kg when Easidew sensor fitted) Vacuum Rating:	Not vacuum rated

Material:	Block & Cap: 316 stainless steel Flow control valve: aluminium Couplings: nickel-plated brass

CUSTOMER SERVICE CONTACT DETAILS

For advice on this, or any other Michell Instruments product, please feel free to contact us on any of the numbers below:

Michell Instruments Ltd.

Unit 9, Nuffield Close, Cambridge, CB4 1SS

Tel: +44(0)1223 434800  Fax: +44(0)1223 434895

E-mail: info@michell.co.uk

Web Site: www.michell-instruments.com

EASIDEW SAMPLER

Users Guide

Michell Instruments GmbH

Industriestrasse 27, 61381 Friedrichsdorf, Deutschland

Tel: +49(0)6172 5917 0  Fax: +49(0)6172 5917 99

E-mail: info@michell.de

Michell Instruments Benelux BV

Kastanjelaan 18b, 2982 CM Ridderkerk, The Netherlands

Tel: +31(0)180 466460  Fax: +31(0)180 415623

E-mail: info@michell-instruments.nl

# Shanghai Jetcon Automation Technology Co.,Ltd

Add: 667 Daning Road Building No.6,Apt.101, Shanghai,China

TEL: 86-021-66527746  86-021-56775537  86-021-28865446

FAX: 86-021-56775537

http: // www.jkong.com

Issue 02, Nov 2003	                   Manual No.97076